
	[image: image1.png]Uniersityof Loadon



                                 
Responding to a HAIER alarm situation 

	

	

	SOP: 2
	Version 1
	Effective Date:
	15th October 2021

	Version number and date:
	15th October 2021 (V1)
	Review Date:
	October 2022

	Author:
	Ash Sameja
	Title:
	Laboratory Manager

	Approved by:
	Penny Lympany
	Date:
	15 /10/2021


Standard Operating Procedure (SOP) Responding to an alarm situation

	Table of Contents
	Page

	1. Purpose
	1

	2.  Definitions/Equipment
	1

	3. Scope
	2

	4. Process
	2

	5. Reference
	6

	6. Contacts
	5


1. Purpose

1.1 Jenner Wing has several cold storage units requiring regular temperature monitoring. Along with the Rees Scientific system (SOP 1), the Haier System is used to monitor some of these units.  This second SOP provides instructions for responding to an alarm.

1.2 Applicable to -80C, -40C, -20C freezers, fridges, Liquid Nitrogen vessels and Cold Rooms fitted with Haier monitoring temperature probes. Staff will be responsible for their own equipment and will need to respond to any alarm situation.
2. Definitions / Equipment
2.1. Haier monitoring system



 Real time is a wireless monitoring system from Haier Biomedical and is used to monitor temperatures of cold storage equipment such as fridges, cold room, -80C freezers & Liquid Nitrogen tanks within Jenner Wing. The system monitors the temperature 24/7 365 days a year and sends alarm notification by phone calls, text messages and email. The added feature of the system is that it allows users to access and interact in real-time with the main control system for their own storage vessels via an external web access. External access is only possible by authenticated staff on the system and requires a user name and password. Please contact Administrators responsible for the Haier monitoring System. 
Head of Laboratory Services - Penny Lympany (plympany@sgul.ac.uk) 
Laboratory Manager - Ash Sameja (asameja@gul.ac.uk) .

Once authorized, the user can access the system using any browser from PC, tablet or Smartphone by typing the following in URL https://www.realtime-online.com/ . A login screen will then appear. 
3. Scope

This SOP covers the process for responding to an alarm triggered by the Haier monitoring system and applies to staff who are on the Haier system telephone and email list. The individual department may also have their own local action policy which they will need to follow on a call out. 
The alarm alert is designed to prevent adverse situations and loss of precious stored research samples. 

4. Process
4.1. Alarm Situations
4.1.1. Each device will have parameters of upper and lower temperature set and if the limits are crossed, an alarm is triggered. If the alarms are activated, the System’s auto dialler will initiate sending calls and text messages to ALL the programmed numbers to say that an alarm has been raised. 
4.1.2. At the same time as the phone calls and text alerts, an email will be sent to the recipients giving details of the alarm situation.
Phone –

If unanswered, check your Voicemail.

SMS Text –

From Realtime online – Haier – St Georges University

[image: image2.png]Realtime Online

Haier — St Georges University
London

J

temperature:-56.7 C
Max:-60 C

Min:-90 C

Aug 18 2021 14:02:59





Email 
with Time stamp, Sensor Input, low and a high temperature limit. The email will be from Invisible-systems - no-reply@realtime-online.com  See example below.
[image: image3.png]The following probes have performed outside the defined parameters:

ftemperature _|J2.115 - 37 6327555 |Aug 18 2021|14:02559 |Alam _|°C





PLEASE NOTE : Please record your actions and any remedial work regarding this notification by logging into your realtime-online web portal. No further warnings will be generated for the above sensor(s) for another 24 hours.
4.2. Viewing probe data.
4.2.1. Open any browser. Type in URL – https://www.realtime-online.com  Log in with your Full Username and password. 
[image: image4.png]Forgotten Pa

This ste uses cookies. By logging in to this site you are agreeing to the
use of cookies
See our Cookie Policy GDPR Privacy Policy





4.2.1 Once logged in you will see the alarm state of the freezer in flashing red.
[image: image5.png]CONDITIONS PROFILE

REPORTS

KLD80-02JW1-144 Today

s o

H

3 -0

2

g -100
13/10/2021  01:00 02:00 03:00 04:00 05:00 06:00
KLD80-3-Biobank KLD1131-01 KLD1143-01 KL

Condition Snapshot - Haier - SGUL - PIDRG
<M de: > Max

Sensor
KLD80-3-Biobank
KLD1131:01
KLD1143-01
KLD20-01JW2-206.
KLDE0-02JW1-144.

17/101/2090

1 aet Mats Dacied 9024.10.12 11





4.3. To cancel and acknowledge the alarm
4.3.1. Double-click on Date.
[image: image6.png]Recent Notifications

2021-10-1311:1920 _|KLD80-3-Biobank _|Condition Alam





You will see details of the alarm

[image: image7.png]Alarm Details

Alarm ID: 11030578

Sensor: KLD80-3-Biobank (9719748)
Date: 2021-10-13

Time: 11:19:20

Reading: -82.3 °C

Min Threshold: 70 °C

Max Threshold: -65 °C

Action

1~ Temperature Excursion

[Save]




To acknowledge the alarm (turns green status from red) and record the actions undertaken, use the “Edit” button. This information entered will be used for audit purposes.
[image: image8.png]Details

Search: Start/End Date:  2021-07-14-2021-10-14  Notification Type:  all VAl v
BreachDate  Email SMS TTS Location Reading Severity Status Actioned  Action Notes User  Options

2021-10-13 Yes  Yes No KLD80-3- -825°C  Alarm 2021-10-14  Set point Testing Ashraf

11:34:14 Biobank 110526 alteredas  alarmfor  Sameja

management PIDRG users

request (06)  and discuss
procedure to
follow.




4.3.2. If you think the problem is a false alarm you can reset and enable.

4.3.3. If the situation requires inspection of the cold storage devices or probes, please contact your ‘on call’ partner to agree the required response – e.g. site visit
4.3.4. If you are arriving out of hours, please inform St. George’s University security on ext. 0909, giving your name and contact numbers so that they know you are on site. Upon arrival check the device. 
4.4. Checklist
4.4.1. Check for power (section 4.5)

4.4.2. Check the appliance. If the door has been left open, close it and monitor the readouts to see if the temperature starts to fall. 
4.4.3. If the door has not been left open, monitor the readouts to see if the temperature is falling or still rising. If the temperature continues to rise, implement the relevant disaster recovery SOP (check with your Manager).
4.5. The electricity and generator have failed
4.5.1. The appropriate fridges and freezers are on a supply connected to a back-up generator which maintains the supply in event of a shutdown. When the generator starts it may trip the alarm. 
4.5.2. If the power has not been restored the University disaster recovery SOPs will be initiated.
        4.6 Backup Emergency Freezer 
If the freezer has failed, then you will need to rescue your samples by removing and placing them to the nearest Backup freezer. If this action is to be taken during Office hours, then please contact the Technical Research Operations Team (ext 0139) for assistance.
5 REFERENCE (See FLOW DIAGRAM on page 6)
CONTACTS

	Penny Lympany
	Head of Laboratory Services
	Ext 1603
	plympany@sgul.ac.uk

	Ash Sameja
	Laboratory Manager
	Ext 2428
	asameja@sgul.ac.uk


Freezer Alarm FLOW DIAGRAM




Alarm Activated





Report to Estates & Facilities


For Electric /Air condition failure








Make your way to University





Check on PC (Haier)


Or


Physically check freezer





Check on-line


Check Graph readings


Check with reception (0909) if there has been power failure





Yes





Transfer samples to Emergency Freezer





2) Check mains power cable plugged in?








No





NO





1) Is LED display on?














Allow 30mins


For freezer to settle





YES





Office Hours





Has someone opened door?





Phone call





Email & Phone Alert





Out of Office Hours








2

